Adresse postale : CS 30206 78961 Saint Quentin en Yvelines Cedex Société anonyme au capital de 38113€ - RCS Versailles B 420 668 402 – Code NAF 7010Z 1/11 Boston Scientific International S.A. EMEA Headquarters Siège social : Parc du Val Saint Quentin – 2 rue René Caudron 78960 Voisins le Bretonneux – France Tel 33 (0)1 39 30 97 00 Fax 33 (0)1 39 30 97 99 www.bostonscientific.com

<Hospital_Name>

<Users_Name>

<Department>

<Customer_Address>

<Zip_Code> <City>

<Country_Name>

# <Reference: 97507023-FA>									19 December 2025

SRN: US-MF-000004702

# Urgent Field Safety Notice - Urgent Medical Device Recall

AXIOS™ Stent and Electrocautery Enhanced Delivery System

Dear Materials Manager and/or Healthcare Provider,

Description of Issue and Clinical Impact:

Boston Scientific is initiating a removal of certain AXIOS™ Stent and Electrocautery Enhanced Delivery Systems (6 mm × 8 mm, 8 mm × 8 mm, and 20 mm × 10 mm), detailed in Attachment 1. This action is being taken due to increased reports of stent deployment and expansion issues with these configurations. This issue only occurs at the time of delivery of the stent and is expected to be noticed by the physician. Successfully implanted stents are not affected by this issue.

The most common outcome reasonably foreseeable resulting from difficulty deploying the stent is prolongation of the procedure to exchange the device for a new one.

The most serious reasonably foreseeable outcome in cases where the first flange is unable to deploy or expand would be a cascade of events resulting in additional intervention, either endoscopic or surgical, to remove the stent and close the puncture site.

The serious injury events also included 3 deaths; these events were associated with use outside of indication (EDGE procedure, gastrojejunostomy) or investigational use (AXIOS GE for GOO IDE study) and required a cascade of events with special circumstances such as decision for palliative care, altered anatomy or acute unrelated procedural complication, all of which limited further treatment options.

Recommendations:

Patients who have been treated with a successfully implanted AXIOS™ stent should continue to follow standard of care, as the issue only occurs at the time of delivery of the stent. Any additional management should be at the discretion of their physician based on individual circumstances.

Our records indicate that your facility has received some of the concerned product. The table below (Attachment 1) provides a complete list of all affected products, including Product Description, Material Number (UPN), GTIN, Lot/Batch numbers and expiry date. Please note that only the devices listed below are affected. No other Boston Scientific product is involved in this Field Safety Notice.

Further distribution or use of any remaining product affected by this action should cease immediately.

INSTRUCTIONS:

1- Please immediately discontinue use of the Boston Scientific product reported in the list and remove all of the affected units from your inventory, regardless of where these units are stored in your facility. Segregate the units in a secure place, pending return to Boston Scientific.

2- Immediately post this information in a visible location near the affected products to ensure this information is readily accessible to all handlers and users of the device.

3- Please complete the attached Verification Form even if you do not have any product to return.

4- When completed, please return the Verification Form to your local Boston Scientific office for the attention of <Customer_Service_Fax_Number> on or before 26 January 2026.

5- If you have products to return, please package them in an appropriate shipping box. After receipt of the Verification Form, Boston Scientific will contact you to arrange return.

6- Please pass this notice to any health professional from your organization that needs to be aware and to any organization where the potentially affected devices have been transferred (if appropriate). Please provide Boston Scientific with details of any affected devices that have been transferred to other organizations (if appropriate).

7- Boston Scientific reminds healthcare professionals that AXIOS Stents should only be used as indicated in Instructions For Use.

Your Competent Authority has been notified of this Field Safety Notice.

Any adverse events or quality concerns associated with use of these devices should be reported to Boston Scientific.

Patient safety is Boston Scientific’s highest priority. As such, we are committed to transparent communication with physicians and healthcare professionals to ensure you have timely, relevant information for managing your patients.If you have any questions or would like assistance with this Field Safety Notice, please contact your local Sales Representative.

2/11

Yours sincerely,

Attachment: Verification Form

FOR BOSTON SCIENTIFIC INTERNAL USE ONLY

Account Email: <Contact Email>

Language: <Language(s)>

LFAC Team: <LFAC_Distribution_Email_Address>

Country Code-Sold to: <Country_Code>-<Sold_to_b>

Attachment 1- Affected Device List

| Product Description | UPN | GTIN | Lot # |  |  |  | Expiration Date Range |
| --- | --- | --- | --- | --- | --- | --- | --- |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553520 | 8714729904540 | 36067957 | 36459567 | 36680231 | 37427362 | September 12, 2026 through April 28, 2027 |
|  |  |  | 36067958 | 36459568 | 36680232 | 37437936 |  |
|  |  |  | 36067959 | 36459569 | 36683260 | 37438283 |  |
|  |  |  | 36067960 | 36461959 | 36683261 | 37438286 |  |
|  |  |  | 36067961 | 36462320 | 36683262 | 37438287 |  |
|  |  |  | 36099410 | 36468392 | 36693971 | 37439978 |  |
|  |  |  | 36102139 | 36468393 | 36693972 | 37440320 |  |
|  |  |  | 36111317 | 36468394 | 36717713 | 37451304 |  |
|  |  |  | 36111318 | 36502795 | 36717714 | 37456594 |  |
|  |  |  | 36111319 | 36506161 | 36717715 | 37457216 |  |
|  |  |  | 36132439 | 36506163 | 36717819 | 37457217 |  |
|  |  |  | 36132800 | 36509326 | 36721863 | 37457269 |  |
|  |  |  | 36136724 | 36520053 | 36731805 | 37459571 |  |
|  |  |  | 36136725 | 36523274 | 36731806 | 37463623 |  |
|  |  |  | 36144054 | 36528114 | 36738965 | 37463624 |  |
|  |  |  | 36154527 | 36536932 | 36738966 | 37463625 |  |
|  |  |  | 36154528 | 36536933 | 36743835 | 37463628 |  |
|  |  |  | 36154529 | 36548030 | 36750922 | 37469756 |  |
|  |  |  | 36156019 | 36548031 | 36750924 | 37469757 |  |
|  |  |  | 36164532 | 36548850 | 36753179 | 37469758 |  |
|  |  |  | 36164990 | 36548851 | 36753400 | 37470669 |  |
|  |  |  | 36176206 | 36552680 | 36781577 | 37472148 |  |
|  |  |  | 36176207 | 36559035 | 36781578 | 37475122 |  |
|  |  |  | 36177448 | 36559036 | 36781579 | 37481983 |  |
|  |  |  | 36191239 | 36567446 | 36784378 | 37486079 |  |
|  |  |  | 36199389 | 36567447 | 36784379 | 37518915 |  |
|  |  |  | 36217761 | 36567448 | 36788332 | 37518916 |  |
|  |  |  | 36241713 | 36574150 | 36788333 | 37518917 |  |
|  |  |  | 36251361 | 36581585 | 36788334 | 37519094 |  |
|  |  |  | 36252913 | 36581586 | 36792650 | 37532420 |  |
|  |  |  | 36271441 | 36582790 | 36799316 | 37532421 |  |
|  |  |  | 36271442 | 36582791 | 36832750 | 37533309 |  |
|  |  |  | 36281910 | 36585188 | 36832837 | 37543372 |  |
|  |  |  | 36281911 | 36585189 | 36832838 | 37558499 |  |
|  |  |  | 36292639 | 36591087 | 36833682 | 37558542 |  |
|  |  |  | 36292681 | 36592499 | 36834573 | 37567252 |  |
|  |  |  | 36300838 | 36592660 | 36843542 | 37567253 |  |
|  |  |  | 36317858 | 36596040 | 36844231 | 37567400 |  |
|  |  |  | 36317859 | 36596041 | 36846161 | 37568848 |  |
|  |  |  | 36318240 | 36609818 | 36846163 | 37568849 3/11 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553520 | 8714729904540 | 36323786 | 36609819 | 36864951 | 37571930 | September 12, 2026 through April 28, 2027 |
|  |  |  | 36323787 | 36610400 | 36864952 | 37685005 |  |
|  |  |  | 36331036 | 36610401 | 36865501 | 37686597 |  |
|  |  |  | 36331038 | 36621583 | 36865502 | 37689957 |  |
|  |  |  | 36397921 | 36621584 | 36865504 | 37689958 |  |
|  |  |  | 36397922 | 36634546 | 36865553 | 37698144 |  |
|  |  |  | 36397923 | 36634547 | 36868065 | 37704842 |  |
|  |  |  | 36421238 | 36664793 | 36916329 | 37761034 |  |
|  |  |  | 36424644 | 36673195 | 36916330 | 37775246 |  |
|  |  |  | 36424645 | 36673557 | 36917403 | 37776579 |  |
|  |  |  | 36427248 | 36673558 | 36917404 | 37776720 |  |
|  |  |  | 36446498 | 36675688 | 36917993 | 37776721 |  |
|  |  |  | 36447200 | 36675689 | 37423232 | 37780826 |  |
|  |  |  | 36447201 | 36675690 | 37423233 | 37785688 |  |
|  |  |  | 36449569 | 36675692 | 37427258 | 37785689 |  |
|  |  |  | 36449570 | 36677844 | 37427259 | 37811064 |  |
|  |  |  | 36449571 | 36677845 | 37427360 | 37817121 |  |
|  |  |  | 36458829 | 36677846 | 37427361 |  |  |
|  | M00553530 | 8714729904557 | 35985705 | 36517139 | 36766371 | 37498192 | September 2, 2026 through April 23, 2027 |
|  |  |  | 35986358 | 36517220 | 36783265 | 37511243 |  |
|  |  |  | 36147098 | 36525208 | 36792467 | 37519095 |  |
|  |  |  | 36147222 | 36525209 | 36792468 | 37519096 |  |
|  |  |  | 36165642 | 36525210 | 36792469 | 37519097 |  |
|  |  |  | 36176423 | 36534514 | 36792470 | 37520927 |  |
|  |  |  | 36190575 | 36534518 | 36792471 | 37530320 |  |
|  |  |  | 36191222 | 36536259 | 36799315 | 37530321 |  |
|  |  |  | 36194448 | 36536480 | 36815090 | 37530333 |  |
|  |  |  | 36218745 | 36536481 | 36815092 | 37533308 |  |
|  |  |  | 36227778 | 36551000 | 36833683 | 37541933 |  |
|  |  |  | 36227779 | 36551001 | 36845422 | 37571927 |  |
|  |  |  | 36241500 | 36642564 | 36851792 | 37571928 |  |
|  |  |  | 36271259 | 36642565 | 36873631 | 37579364 |  |
|  |  |  | 36280398 | 36642756 | 36873632 | 37579366 |  |
|  |  |  | 36293880 | 36645775 | 36873906 | 37579367 |  |
|  |  |  | 36300839 | 36663654 | 36904228 | 37580630 |  |
|  |  |  | 36301620 | 36677842 | 36904230 | 37585525 |  |
|  |  |  | 36343103 | 36677843 | 36907002 | 37585526 |  |
|  |  |  | 36343104 | 36682759 | 36907004 | 37589261 |  |
|  |  |  | 36364646 | 36684887 | 36907028 | 37589262 |  |
|  |  |  | 36425887 | 36684888 | 37102533 | 37594668 |  |
|  |  |  | 36469811 | 36730665 | 37102534 | 37649481 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553530 | 8714729904557 | 36471228 | 36730666 | 37102535 | 37649482 4/11 | September 2, 2026 through April 23, 2027 |
|  |  |  | 36491879 | 36747824 | 37102536 | 37649483 |  |
|  |  |  | 36492660 | 36747825 | 37102537 | 37650631 |  |
|  |  |  | 36506029 | 36753898 | 37459567 | 37694052 |  |
|  |  |  | 36510220 | 36753899 | 37459568 | 37699011 |  |
|  |  |  | 36510221 | 36757493 | 37459569 | 37699012 |  |
|  |  |  | 36513401 | 36757494 | 37459570 | 37708202 |  |
|  |  |  | 36513402 | 36761827 | 37463133 | 37779125 |  |
|  |  |  | 36513403 | 36766370 | 37470670 |  |  |
|  | M00553560 | 8714729951100 | 35793920 | 36310764 | 36683263 | 37263554 | February 6, 2026 through October 28, 2026 |
|  |  |  | 35952745 | 36310765 | 36684615 | 37263555 |  |
|  |  |  | 35961004 | 36310766 | 36685905 | 37268080 |  |
|  |  |  | 35995006 | 36310767 | 36685906 | 37268081 |  |
|  |  |  | 35995007 | 36310768 | 36804376 | 37268098 |  |
|  |  |  | 35995165 | 36310769 | 36814735 | 37269066 |  |
|  |  |  | 36001950 | 36447543 | 36816919 | 37269068 |  |
|  |  |  | 36001951 | 36447544 | 36818381 | 37273840 |  |
|  |  |  | 36001952 | 36447545 | 36818382 | 37273841 |  |
|  |  |  | 36001953 | 36447546 | 36821877 | 37274479 |  |
|  |  |  | 36002902 | 36449567 | 36823193 | 37274863 |  |
|  |  |  | 36002903 | 36449568 | 36823194 | 37275620 |  |
|  |  |  | 36002904 | 36458833 | 36828324 | 37276239 |  |
|  |  |  | 36003946 | 36458834 | 36828325 | 37276361 |  |
|  |  |  | 36003947 | 36461125 | 36829194 | 37276363 |  |
|  |  |  | 36003948 | 36478581 | 36829196 | 37276364 |  |
|  |  |  | 36012845 | 36478582 | 36835382 | 37308911 |  |
|  |  |  | 36012846 | 36478583 | 36835383 | 37323226 |  |
|  |  |  | 36012847 | 36478584 | 36838544 | 37323227 |  |
|  |  |  | 36012848 | 36478585 | 36838545 | 37323228 |  |
|  |  |  | 36015761 | 36481333 | 36838546 | 37323229 |  |
|  |  |  | 36021049 | 36481334 | 36838547 | 37324379 |  |
|  |  |  | 36021050 | 36487762 | 36839663 | 37325180 |  |
|  |  |  | 36021051 | 36487763 | 36839664 | 37325181 |  |
|  |  |  | 36021052 | 36487764 | 36847979 | 37325182 |  |
|  |  |  | 36027054 | 36487765 | 36848665 | 37325354 |  |
|  |  |  | 36027055 | 36493311 | 36848666 | 37331113 |  |
|  |  |  | 36027056 | 36493313 | 36851793 | 37379513 |  |
|  |  |  | 36027057 | 36493314 | 36851794 | 37380940 |  |
|  |  |  | 36027058 | 36494812 | 36853105 | 37380941 |  |
|  |  |  | 36027059 | 36494813 | 36853106 | 37380942 |  |
|  |  |  | 36027060 | 36503314 | 36853594 | 37382298 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553560 | 8714729951100 | 36027061 | 36503315 | 36853595 | 37382299 | February 6, 2026 through October 28, 2026 |
|  |  |  | 36038410 | 36503316 | 36854299 | 37384504 |  |
|  |  |  | 36038411 | 36503317 | 36856185 | 37384505 |  |
|  |  |  | 36040000 | 36503459 | 36856186 | 37386277 5/11 |  |
|  |  |  | 36040001 | 36504640 | 36856187 | 37386278 |  |
|  |  |  | 36040002 | 36504702 | 36856188 | 37443960 |  |
|  |  |  | 36040003 | 36509327 | 36856340 | 37443961 |  |
|  |  |  | 36040004 | 36517739 | 36856341 | 37443962 |  |
|  |  |  | 36048795 | 36517821 | 36856342 | 37451306 |  |
|  |  |  | 36048796 | 36517822 | 36866851 | 37451307 |  |
|  |  |  | 36048797 | 36519644 | 36866852 | 37456598 |  |
|  |  |  | 36048798 | 36533327 | 36866853 | 37457260 |  |
|  |  |  | 36049282 | 36533328 | 36868980 | 37460647 |  |
|  |  |  | 36049283 | 36548037 | 36868981 | 37460648 |  |
|  |  |  | 36049284 | 36548038 | 36868982 | 37461035 |  |
|  |  |  | 36050387 | 36548039 | 36868983 | 37461036 |  |
|  |  |  | 36050388 | 36548060 | 36868984 | 37463128 |  |
|  |  |  | 36050389 | 36548061 | 36869884 | 37463129 |  |
|  |  |  | 36059222 | 36551002 | 36869885 | 37501715 |  |
|  |  |  | 36059223 | 36551003 | 36869886 | 37501716 |  |
|  |  |  | 36059224 | 36551004 | 36869887 | 37501717 |  |
|  |  |  | 36059225 | 36559039 | 36869888 | 37510239 |  |
|  |  |  | 36062077 | 36559580 | 36869889 | 37511240 |  |
|  |  |  | 36062078 | 36559581 | 36869890 | 37514184 |  |
|  |  |  | 36062079 | 36559582 | 36869891 | 37514185 |  |
|  |  |  | 36064449 | 36559583 | 36875736 | 37514972 |  |
|  |  |  | 36065116 | 36560842 | 36875737 | 37524755 |  |
|  |  |  | 36065117 | 36560843 | 36877081 | 37524756 |  |
|  |  |  | 36067962 | 36562844 | 36877082 | 37524758 |  |
|  |  |  | 36067963 | 36562845 | 36877084 | 37530261 |  |
|  |  |  | 36067964 | 36565931 | 36877085 | 37536574 |  |
|  |  |  | 36069958 | 36566528 | 36880763 | 37538606 |  |
|  |  |  | 36069959 | 36567449 | 36880764 | 37541652 |  |
|  |  |  | 36070600 | 36570785 | 36897772 | 37543964 |  |
|  |  |  | 36070608 | 36570786 | 36897773 | 37543965 |  |
|  |  |  | 36070609 | 36570787 | 36959806 | 37543966 |  |
|  |  |  | 36070610 | 36572359 | 36959807 | 37544646 |  |
|  |  |  | 36070611 | 36581587 | 37004726 | 37544647 |  |
|  |  |  | 36072736 | 36583578 | 37004727 | 37544648 |  |
|  |  |  | 36072737 | 36583579 | 37059011 | 37548357 |  |
|  |  |  | 36092333 | 36591069 | 37060652 | 37548358 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553560 | 8714729951100 | 36092334 | 36591070 | 37067633 | 37549491 | February 6, 2026 through October 28, 2026 |
|  |  |  | 36092335 | 36591071 | 37069356 | 37549492 |  |
|  |  |  | 36099411 | 36591072 | 37069357 | 37552245 |  |
|  |  |  | 36099412 | 36612260 | 37075264 | 37552246 |  |
|  |  |  | 36099413 | 36612261 | 37075265 | 37552247 |  |
|  |  |  | 36101743 | 36616444 | 37078233 | 37552248 |  |
|  |  |  | 36101744 | 36616445 | 37078234 | 37554898 |  |
|  |  |  | 36101745 | 36619789 | 37093209 | 37567401 |  |
|  |  |  | 36101746 | 36619790 | 37093210 | 37567402 |  |
|  |  |  | 36101747 | 36619791 | 37093302 | 37567403 6/11 |  |
|  |  |  | 36102812 | 36619792 | 37093303 | 37567404 |  |
|  |  |  | 36102813 | 36619793 | 37096513 | 37567405 |  |
|  |  |  | 36102814 | 36621321 | 37096514 | 37567406 |  |
|  |  |  | 36135590 | 36621322 | 37099739 | 37601264 |  |
|  |  |  | 36135591 | 36621323 | 37132500 | 37601265 |  |
|  |  |  | 36144122 | 36625032 | 37132501 | 37601266 |  |
|  |  |  | 36144123 | 36625033 | 37132502 | 37611305 |  |
|  |  |  | 36147096 | 36625034 | 37134087 | 37611306 |  |
|  |  |  | 36147097 | 36625035 | 37137422 | 37611308 |  |
|  |  |  | 36154534 | 36626350 | 37137423 | 37617701 |  |
|  |  |  | 36154536 | 36626351 | 37140068 | 37617702 |  |
|  |  |  | 36154537 | 36626352 | 37152588 | 37625483 |  |
|  |  |  | 36154538 | 36633245 | 37153325 | 37625484 |  |
|  |  |  | 36155116 | 36633246 | 37157514 | 37644410 |  |
|  |  |  | 36155117 | 36640831 | 37161889 | 37649490 |  |
|  |  |  | 36156264 | 36641049 | 37161890 | 37651467 |  |
|  |  |  | 36156265 | 36641909 | 37164067 | 37655139 |  |
|  |  |  | 36164534 | 36641910 | 37164068 | 37655281 |  |
|  |  |  | 36164535 | 36642073 | 37167102 | 37660774 |  |
|  |  |  | 36164536 | 36642074 | 37167103 | 37660775 |  |
|  |  |  | 36165655 | 36642075 | 37167775 | 37660777 |  |
|  |  |  | 36165656 | 36642562 | 37176653 | 37660778 |  |
|  |  |  | 36165657 | 36643731 | 37176657 | 37665497 |  |
|  |  |  | 36165658 | 36643732 | 37194095 | 37667998 |  |
|  |  |  | 36197586 | 36645532 | 37204177 | 37672121 |  |
|  |  |  | 36197587 | 36645533 | 37204178 | 37675664 |  |
|  |  |  | 36197588 | 36645534 | 37204179 | 37675665 |  |
|  |  |  | 36197589 | 36647015 | 37204240 | 37689944 |  |
|  |  |  | 36199391 | 36662774 | 37204241 | 37689945 |  |
|  |  |  | 36200602 | 36662775 | 37204809 | 37696688 |  |
|  |  |  | 36200603 | 36662776 | 37204810 | 37720785 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553560 | 8714729951100 | 36207756 | 36662778 | 37204811 | 37723579 | February 6, 2026 through October 28, 2026 |
|  |  |  | 36207757 | 36662779 | 37207470 | 37731195 |  |
|  |  |  | 36207758 | 36664061 | 37207471 | 37731196 |  |
|  |  |  | 36209378 | 36664062 | 37208948 | 37731197 |  |
|  |  |  | 36209379 | 36664063 | 37217600 | 37742018 |  |
|  |  |  | 36218689 | 36664064 | 37217605 | 37745174 |  |
|  |  |  | 36218690 | 36664065 | 37217607 | 37746915 |  |
|  |  |  | 36218746 | 36664567 | 37240010 | 37750340 |  |
|  |  |  | 36218747 | 36664568 | 37240011 | 37750341 |  |
|  |  |  | 36219228 | 36664794 | 37252330 | 37754384 |  |
|  |  |  | 36271447 | 36664796 | 37252331 | 37754385 |  |
|  |  |  | 36271448 | 36665694 | 37252333 | 37771353 |  |
|  |  |  | 36271456 | 36665695 | 37252334 | 37771355 |  |
|  |  |  | 36271457 | 36669412 | 37252335 | 37777208 7/11 |  |
|  |  |  | 36271458 | 36669413 | 37257420 | 37778063 |  |
|  |  |  | 36271459 | 36669414 | 37257421 | 37780175 |  |
|  |  |  | 36271840 | 36669415 | 37257422 | 37780176 |  |
|  |  |  | 36303991 | 36669416 | 37257423 | 37809877 |  |
|  |  |  | 36303992 | 36669417 | 37261686 | 37809878 |  |
|  |  |  | 36303993 | 36669418 | 37262942 | 37809879 |  |
|  |  |  | 36310699 | 36674810 | 37262943 | 37811060 |  |
|  |  |  | 36310760 | 36674811 | 37262949 | 37811061 |  |
|  |  |  | 36310761 | 36674812 | 37262950 | 37811062 |  |
|  |  |  | 36310762 | 36674813 | 37263184 | 37811063 |  |
|  |  |  | 36310763 | 36677847 | 37263185 | 37817124 |  |
|  | M00553660 | 8714729951179 | 35961006 | 36359567 | 36895363 | 37340131 | February 28, 2026 through October 29, 2026 |
|  |  |  | 35961007 | 36362263 | 36895364 | 37340270 |  |
|  |  |  | 35962524 | 36362264 | 36958216 | 37340271 |  |
|  |  |  | 35967379 | 36362265 | 36958217 | 37341078 |  |
|  |  |  | 35968520 | 36363150 | 36962750 | 37341079 |  |
|  |  |  | 35968521 | 36363151 | 36970218 | 37344296 |  |
|  |  |  | 35968597 | 36364647 | 36970219 | 37344297 |  |
|  |  |  | 35968598 | 36372361 | 36980310 | 37346443 |  |
|  |  |  | 35971077 | 36372362 | 36981856 | 37346445 |  |
|  |  |  | 35971078 | 36377757 | 36985933 | 37348330 |  |
|  |  |  | 35971079 | 36378548 | 36988632 | 37348331 |  |
|  |  |  | 35972100 | 36380549 | 36990382 | 37352140 |  |
|  |  |  | 35982272 | 36388531 | 36990383 | 37352141 |  |
|  |  |  | 35982273 | 36390612 | 36991168 | 37354943 |  |
|  |  |  | 35982274 | 36396985 | 36991170 | 37354944 |  |
|  |  |  | 35982275 | 36399492 | 36991683 | 37357958 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553660 | 8714729951179 | 35982276 | 36403441 | 36992721 | 37357959 | February 28, 2026 through October 29, 2026 |
|  |  |  | 35982277 | 36408983 | 37014102 | 37358120 |  |
|  |  |  | 35982278 | 36408984 | 37014103 | 37359884 |  |
|  |  |  | 35982279 | 36417820 | 37014104 | 37359885 |  |
|  |  |  | 35982420 | 36417821 | 37027339 | 37365502 |  |
|  |  |  | 35982424 | 36418721 | 37028180 | 37365503 |  |
|  |  |  | 35982425 | 36418722 | 37028181 | 37365504 |  |
|  |  |  | 35985313 | 36424650 | 37030756 | 37365505 |  |
|  |  |  | 35985314 | 36424651 | 37032104 | 37375861 |  |
|  |  |  | 35993619 | 36424652 | 37036344 | 37388591 |  |
|  |  |  | 35993720 | 36425889 | 37036345 | 37388592 |  |
|  |  |  | 35993721 | 36425890 | 37037505 | 37388593 |  |
|  |  |  | 36109761 | 36425891 | 37037506 | 37400608 |  |
|  |  |  | 36109763 | 36425892 | 37040169 | 37402464 |  |
|  |  |  | 36109764 | 36462323 | 37041012 | 37402466 |  |
|  |  |  | 36109765 | 36462324 | 37041435 | 37410374 |  |
|  |  |  | 36109766 | 36469427 | 37041436 | 37411886 |  |
|  |  |  | 36111314 | 36469428 | 37046449 | 37411887 |  |
|  |  |  | 36111315 | 36469812 | 37046450 | 37415923 8/11 |  |
|  |  |  | 36118993 | 36472250 | 37046928 | 37415924 |  |
|  |  |  | 36118994 | 36472253 | 37046929 | 37415925 |  |
|  |  |  | 36122066 | 36472254 | 37048665 | 37427250 |  |
|  |  |  | 36122067 | 36475370 | 37049464 | 37429329 |  |
|  |  |  | 36122075 | 36478478 | 37049465 | 37429330 |  |
|  |  |  | 36132801 | 36478479 | 37049874 | 37438134 |  |
|  |  |  | 36132802 | 36572740 | 37050321 | 37438135 |  |
|  |  |  | 36132803 | 36572741 | 37050322 | 37439609 |  |
|  |  |  | 36132804 | 36573582 | 37050323 | 37439610 |  |
|  |  |  | 36134972 | 36573583 | 37055465 | 37439611 |  |
|  |  |  | 36144058 | 36573587 | 37055466 | 37451308 |  |
|  |  |  | 36144059 | 36573588 | 37055467 | 37451309 |  |
|  |  |  | 36144120 | 36573589 | 37056013 | 37454250 |  |
|  |  |  | 36144121 | 36573590 | 37056014 | 37454251 |  |
|  |  |  | 36174475 | 36591073 | 37057032 | 37456595 |  |
|  |  |  | 36174476 | 36591074 | 37059006 | 37456596 |  |
|  |  |  | 36178005 | 36591075 | 37059010 | 37459573 |  |
|  |  |  | 36187588 | 36603384 | 37094844 | 37567407 |  |
|  |  |  | 36187589 | 36603385 | 37106164 | 37573089 |  |
|  |  |  | 36187687 | 36603386 | 37106583 | 37573090 |  |
|  |  |  | 36187688 | 36604188 | 37131143 | 37573091 |  |
|  |  |  | 36188887 | 36604189 | 37131144 | 37575640 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553660 | 8714729951179 | 36188888 | 36604190 | 37131145 | 37578876 | February 28, 2026 through October 29, 2026 |
|  |  |  | 36188889 | 36604191 | 37131146 | 37578877 |  |
|  |  |  | 36194446 | 36604748 | 37177253 | 37578879 |  |
|  |  |  | 36194447 | 36604749 | 37177271 | 37579368 |  |
|  |  |  | 36197239 | 36607249 | 37177272 | 37579369 |  |
|  |  |  | 36198580 | 36607250 | 37179955 | 37579370 |  |
|  |  |  | 36217768 | 36607251 | 37179956 | 37582027 |  |
|  |  |  | 36217769 | 36607252 | 37180902 | 37585524 |  |
|  |  |  | 36217770 | 36607253 | 37180903 | 37588798 |  |
|  |  |  | 36227341 | 36610818 | 37182323 | 37588799 |  |
|  |  |  | 36227342 | 36610819 | 37183490 | 37589260 |  |
|  |  |  | 36227343 | 36611060 | 37188257 | 37592382 |  |
|  |  |  | 36228603 | 36615703 | 37188258 | 37592383 |  |
|  |  |  | 36228604 | 36615704 | 37197563 | 37596205 |  |
|  |  |  | 36228605 | 36621589 | 37197578 | 37596206 |  |
|  |  |  | 36235934 | 36621590 | 37198908 | 37596207 |  |
|  |  |  | 36236151 | 36623324 | 37199019 | 37596208 |  |
|  |  |  | 36240589 | 36623325 | 37199360 | 37601263 |  |
|  |  |  | 36240590 | 36623326 | 37199361 | 37602176 |  |
|  |  |  | 36241715 | 36693849 | 37199972 | 37602177 |  |
|  |  |  | 36241716 | 36693850 | 37199973 | 37602178 |  |
|  |  |  | 36241717 | 36693851 | 37209169 | 37602179 |  |
|  |  |  | 36251363 | 36693852 | 37209170 | 37603348 |  |
|  |  |  | 36251364 | 36694699 | 37209171 | 37607457 |  |
|  |  |  | 36253000 | 36695020 | 37211282 | 37607459 9/11 |  |
|  |  |  | 36253008 | 36695021 | 37211285 | 37607540 |  |
|  |  |  | 36253041 | 36695023 | 37212243 | 37625485 |  |
|  |  |  | 36260860 | 36698247 | 37212244 | 37625486 |  |
|  |  |  | 36260861 | 36698249 | 37212245 | 37626455 |  |
|  |  |  | 36260862 | 36698251 | 37212246 | 37626456 |  |
|  |  |  | 36261599 | 36698527 | 37213246 | 37631100 |  |
|  |  |  | 36261780 | 36698528 | 37213247 | 37631101 |  |
|  |  |  | 36261781 | 36698529 | 37213248 | 37633190 |  |
|  |  |  | 36261834 | 36698530 | 37213249 | 37633191 |  |
|  |  |  | 36261835 | 36698531 | 37240012 | 37638120 |  |
|  |  |  | 36280395 | 36698532 | 37240013 | 37638121 |  |
|  |  |  | 36280396 | 36706712 | 37240236 | 37638122 |  |
|  |  |  | 36280397 | 36706713 | 37240237 | 37639137 |  |
|  |  |  | 36281912 | 36706714 | 37240249 | 37639138 |  |
|  |  |  | 36281913 | 36706715 | 37240250 | 37644407 |  |
|  |  |  | 36281914 | 36709766 | 37240252 | 37644408 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553660 | 8714729951179 | 36283860 | 36709767 | 37240254 | 37649491 | February 28, 2026 through October 29, 2026 |
|  |  |  | 36292682 | 36709768 | 37242247 | 37649492 |  |
|  |  |  | 36292683 | 36709769 | 37242248 | 37686664 |  |
|  |  |  | 36292684 | 36713255 | 37242249 | 37686665 |  |
|  |  |  | 36292689 | 36713256 | 37243754 | 37686666 |  |
|  |  |  | 36292690 | 36713257 | 37243755 | 37700937 |  |
|  |  |  | 36292691 | 36713258 | 37243756 | 37700938 |  |
|  |  |  | 36294495 | 36717717 | 37247304 | 37706532 |  |
|  |  |  | 36301621 | 36717718 | 37247305 | 37706533 |  |
|  |  |  | 36301622 | 36717719 | 37247306 | 37706534 |  |
|  |  |  | 36301623 | 36717800 | 37283604 | 37708203 |  |
|  |  |  | 36301624 | 36717801 | 37283605 | 37708204 |  |
|  |  |  | 36302290 | 36717802 | 37291851 | 37716332 |  |
|  |  |  | 36302291 | 36718966 | 37291852 | 37716333 |  |
|  |  |  | 36302292 | 36718967 | 37292951 | 37718549 |  |
|  |  |  | 36319024 | 36723909 | 37292952 | 37720553 |  |
|  |  |  | 36319025 | 36723910 | 37296017 | 37720554 |  |
|  |  |  | 36319026 | 36728154 | 37296019 | 37720555 |  |
|  |  |  | 36320016 | 36764237 | 37297941 | 37727831 |  |
|  |  |  | 36320017 | 36764238 | 37300544 | 37730055 |  |
|  |  |  | 36323789 | 36767449 | 37300545 | 37754383 |  |
|  |  |  | 36323790 | 36767450 | 37300546 | 37758168 |  |
|  |  |  | 36324720 | 36771884 | 37302814 | 37758169 |  |
|  |  |  | 36325159 | 36771885 | 37308038 | 37758170 |  |
|  |  |  | 36325560 | 36771886 | 37308039 | 37759554 |  |
|  |  |  | 36325568 | 36785593 | 37325355 | 37761362 |  |
|  |  |  | 36325569 | 36788793 | 37328835 | 37761363 |  |
|  |  |  | 36332101 | 36788794 | 37328836 | 37785682 10/11 |  |
|  |  |  | 36332102 | 36788795 | 37328838 | 37785683 |  |
|  |  |  | 36341577 | 36795369 | 37331112 | 37787089 |  |
|  |  |  | 36342114 | 36795370 | 37332012 | 37787090 |  |
|  |  |  | 36342418 | 36796430 | 37333358 | 37789164 |  |
|  |  |  | 36352008 | 36796431 | 37333362 | 37789168 |  |
|  |  |  | 36352009 | 36804013 | 37337364 | 37789169 |  |
|  |  |  | 36355049 | 36804375 | 37340128 | 37820578 |  |
|  |  |  | 36355050 | 36839665 | 37340129 |  |  |
|  |  |  | 36355752 | 36895362 | 37340130 |  |  |
|  | M00553680 | 191506008086 | 35950784 | 36408034 | 36536257 | 37429734 | August 28, 2026 through April 23, 2027 |
|  |  |  | 35960771 | 36421237 | 36536258 | 37429735 |  |
|  |  |  | 35960772 | 36447549 | 36570788 | 37433804 |  |
|  |  |  | 36156260 | 36447550 | 36570789 | 37433805 |  |
| AXIOS™ Stent and Electrocautery Enhanced Delivery System | M00553680 | 191506008086 | 36156261 | 36520894 | 36578110 | 37451303 | August 28, 2026 through April 23, 2027 |
|  |  |  | 36165653 | 36530482 | 36581584 | 37456593 |  |
|  |  |  | 36167912 | 36530483 | 36603297 | 37495326 |  |
|  |  |  | 36241714 | 36530484 | 36611362 | 37519092 |  |
|  |  |  | 36401105 | 36533326 | 37014108 | 37780825 |  |
|  | M00553690 | 191506008093 | 35960775 | 36303192 | 36576474 | 36907032 | August 29, 2026 through April 22, 2027 |
|  |  |  | 35960777 | 36420219 | 36584299 | 37217612 |  |
|  |  |  | 35962525 | 36425888 | 36604665 | 37217613 |  |
|  |  |  | 36100561 | 36449561 | 36731807 | 37217614 |  |
|  |  |  | 36101748 | 36458830 | 36737997 | 37495327 |  |
|  |  |  | 36111316 | 36461449 | 36743059 | 37511242 |  |
|  |  |  | 36154532 | 36478580 | 36743200 | 37602825 |  |
|  |  |  | 36154533 | 36492661 | 36904231 | 37602826 |  |
|  |  |  | 36241501 | 36495673 | 36904232 | 37694053 |  |
|  |  |  | 36251362 | 36573597 | 36904233 | 37694054 |  |
|  |  |  | 36253016 | 36573598 | 36906587 | 37764852 |  |
|  |  |  | 36293839 | 36573599 | 36907000 |  |  |
|  |  |  | 36301625 | 36574152 | 36907001 |  |  |

11/11

|  | # <Page_No> # <Sold_to> - <Hospital_Name> - <City> - <Country_Name> |
| --- | --- |

# Please Complete the form even if you do not have any affected product & send it to your Local Office: <Customer_Service_Fax_Number>

## ------------------------------------------------------------------------------------------------------------------------------------------------------------------------------

## Verification Form – Urgent Medical Device Recall

AXIOS™ Stent and Electrocautery Enhanced Delivery System

## 97507023-FA

## ------------------------------------------------------------------------------------------------------------------------------------------------------------------------------

1.  We acknowledge receipt of the Boston Scientific Field Safety Notice dated 19 December 2025.

2.  Boston Scientific records indicate you have received the following affected product (additionally please check inventory against complete list of affected product provided)

| Material N° (UPN) | Lot / Batch N° | Customer PO | Qty Sent | Qty to return (Units) |
| --- | --- | --- | --- | --- |

3.  We confirm that all areas where affected product could be located have been checked.

4.  Tick one of these statements*, sign this Form and send it to <Customer_Service_Fax_Number>

- We do not have any affected product.

- We have found affected product(s): Please confirm the quantity to return above. If you are returning product not listed above, please add the UPN, Lot/Batch/Serial number and the quantity to return.

To Return Products:

1.  After receipt of the Verification Form, Boston Scientific will contact you to arrange return.

2.  Prepare the package

3.  Follow the instructions given by your Local Office about collection of the package

Name* __________________________________________________Title __________________________________

Telephone __________________________________________ Email _______________________________________

Customer’ Signature*_________________________________________________ Date*______________________

* Required field									 	      dd/mm/yyyy